HumanWare Accessibility Conformance Report: explorē 12

Revised Section 508 Edition

(Based on VPAT® Version 2.5Rev)

| Field | Report information |
| --- | --- |
| Name of Product/Version | explorē 12 Kit with folding stand; software version 1.1.6.450 |
| Report Date | May 2026 |
| Product Description | Portable electronic video magnifier with 12-inch touchscreen, near and far cameras, magnification, contrast modes, line/blinds, image capture/gallery, USB image export, external display output, folding reading stand, and adjustable settings for low vision use. |
| Contact Information | HumanWare Customer Support: support@humanware.com |
| Notes | This report covers the explorē 12 Kit with folding stand, including the explorē 12 electronic magnifier, folding reading stand, USB charger/adapters, USB cable, high-resolution video cable, carry items, user guide, quick start, related support materials, and HumanWare-produced support videos linked from the HumanWare support page to YouTube. Tablet-only and stand-only variants derive from this kit. explorē 12 is intended for people with limited but usable vision. It is not intended as a blindness product and does not provide speech output or braille output. This report does not evaluate the full HumanWare website or the YouTube player interface. |
| Evaluation Methods Used | Review of the Revised Section 508 VPAT 2.5Rev template, current explorē 12 user guide, technical specifications, product team responses, targeted hands-on testing/confirmation of hardware and embedded software behavior, review of support documentation accessibility, and review of support-service claims. Testing was performed by product personnel with knowledge of common product uses and flows. |

# Applicable Standards/Guidelines

This report covers the degree of conformance for the following accessibility standards/guidelines.

| Standard/Guideline | Included In Report |
| --- | --- |
| Web Content Accessibility Guidelines 2.0 | Level A: Yes Level AA: Yes Level AAA: No |
| Revised Section 508 standards published January 18, 2017 and corrected January 22, 2018 | Yes |

# Terms

The terms used in the Conformance Level information are defined as follows:

Supports: The functionality of the product has at least one method that meets the criterion without known defects or meets with equivalent facilitation.

Partially Supports: Some functionality of the product does not meet the criterion.

Does Not Support: The majority of product functionality does not meet the criterion.

Not Applicable: The criterion is not relevant to the product.

Not Evaluated: The product has not been evaluated against the criterion. This can only be used in WCAG Level AAA criteria.

# WCAG 2.0 Report

Tables 1 and 2 also document conformance with Revised Section 508 Chapter 5 - Software and Chapter 6 - Support Documentation and Services, where applicable.

Scope for WCAG tables: embedded non-web software, electronic support documentation, and HumanWare-produced support videos linked from the support page are included in this report scope and are evaluated based on available captions, transcripts, voice-over, and video content. Web content and authoring tools are treated as Not Applicable unless specifically noted. The full HumanWare website and YouTube player interface are outside this report scope.

## Table 1: Success Criteria, Level A

Notes: Applies to embedded software, electronic support documentation, and included support videos as described in the scope above.

| Criteria | Conformance Level | Remarks and Explanations |
| --- | --- | --- |
| 1.1.1 Non-text Content | Partially Supports | Software: on-device icons and controls are visually presented; core functions are also available through physical buttons and visual labels. There is no screen reader or text-alternative exposure through platform accessibility services. Electronic support documents have alt text for images, diagrams, and graphics. |
| 1.2.1 Audio-only and Video-only (Prerecorded) | Not Applicable | The on-device software does not include prerecorded audio-only or video-only media. Support videos included in this report are prerecorded synchronized media, not audio-only or video-only content. |
| 1.2.2 Captions (Prerecorded) | Supports | The on-device software does not include prerecorded synchronized media. HumanWare-produced support videos included in this report scope are linked from the HumanWare support page to YouTube and include captions. |
| 1.2.3 Audio Description or Media Alternative (Prerecorded) | Partially Supports | The on-device software does not include prerecorded video content. HumanWare-produced support videos include voice-over, captions, and full transcripts. The videos do not include separate audio description tracks. The voice-over describes the demonstrated product tasks, which may provide sufficient context for many videos, but separate audio description has not been provided. |
| 1.3.1 Info and Relationships | Partially Supports | Software: menus and controls use a simple, consistent visual structure, but accessibility semantics are not exposed to assistive technology. Electronic support documents are accessible. |
| 1.3.2 Meaningful Sequence | Supports | Software: primary workflows follow a logical visual and operational sequence. Electronic support documents are accessible and maintain a meaningful reading order. |
| 1.3.3 Sensory Characteristics | Supports | Instructions identify controls by name and location, not only by color, shape, or position. Color is used, but not as the only means of identifying core controls. |
| 1.4.1 Use of Color | Supports | Color is not the only means of access. The product provides named controls, physical button locations, icons, and multiple contrast modes. |
| 1.4.2 Audio Control | Not Applicable | The product does not play audio automatically. |
| 2.1.1 Keyboard | Partially Supports | The product is not designed for a standard external keyboard interface. Core product functions have physical-button equivalents. Selecting items in settings menus requires touch and does not have a physical-button-only selection path. |
| 2.1.2 No Keyboard Trap | Supports | No standard keyboard interface is provided. Physical-button and touch operation do not trap the user in a control or screen by design. |
| 2.2.1 Timing Adjustable | Supports | The auto-off timer is adjustable to 1 minute, 2 minutes, 5 minutes, 10 minutes, or Never. |
| 2.2.2 Pause, Stop, Hide | Supports | The device does not present moving, blinking, scrolling, or auto-updating interface content that prevents use. The live camera view is the core function of the product. |
| 2.3.1 Three Flashes or Below Threshold | Supports | No flashing content occurs during normal operation. |
| 2.4.1 Bypass Blocks | Not Applicable | This criterion does not apply to non-web software or non-web documents in this report scope. |
| 2.4.2 Page Titled | Supports | Software screens and menus use clear labels or context. Electronic support documents include appropriate titles. |
| 2.4.3 Focus Order | Partially Supports | Menu navigation follows a logical visual order. The software does not expose a programmatic focus order to assistive technology. |
| 2.4.4 Link Purpose (In Context) | Supports | On-device software does not rely on links. Electronic support documents use meaningful links. |
| 3.1.1 Language of Page | Supports | The product supports multiple interface languages. The user guide includes English and French content. |
| 3.2.1 On Focus | Supports | No change of context is expected merely from focus or highlighting. Selection or activation is user initiated. |
| 3.2.2 On Input | Supports | Changing settings is user initiated and predictable. |
| 3.3.1 Error Identification | Partially Supports | Some error or status information is visual only. The user guide includes troubleshooting guidance, but the device does not provide nonvisual error identification. |
| 3.3.2 Labels or Instructions | Supports | Controls, menu items, and settings are labeled in the product interface and described in the user guide. |
| 4.1.1 Parsing | Supports | The VPAT 2.5Rev template notes this criterion is always answered Supports for WCAG 2.0 and 508 standards after the September 2023 errata. |
| 4.1.2 Name, Role, Value | Partially Supports | The embedded software does not expose name, role, and value information through platform accessibility services. Core functions are presented visually and through physical controls. |

## Table 2: Success Criteria, Level AA

Notes: Applies to embedded software, electronic support documentation, and included support videos as described in the scope above. WCAG Level AAA is not included in this report.

| Criteria | Conformance Level | Remarks and Explanations |
| --- | --- | --- |
| 1.2.4 Captions (Live) | Not Applicable | The on-device software does not include live synchronized media. |
| 1.2.5 Audio Description (Prerecorded) | Partially Supports | The on-device software does not include prerecorded video content. HumanWare-produced support videos include voice-over, captions, and full transcripts. The videos do not include separate audio description tracks. The voice-over describes the demonstrated product tasks, which may provide sufficient context for many videos, but separate audio description has not been provided. |
| 1.4.3 Contrast (Minimum) | Supports | The product is designed for low vision use and provides multiple high-contrast color combinations, menu color options, screen brightness control, and menu text size options. Contrast ratios have been confirmed to meet requirements. |
| 1.4.4 Resize Text | Supports | The product provides magnification and menu text size settings: Small, Medium, and Large. Live image magnification is available. |
| 1.4.5 Images of Text | Partially Supports | The device is a camera-based magnifier and displays images of text as part of its core function. UI labels are identified. Support documentation provides alt text for images and graphics. |
| 2.4.5 Multiple Ways | Not Applicable | This criterion does not apply to non-web software or non-web documents in this report scope. |
| 2.4.6 Headings and Labels | Supports | Menu items, controls, and documentation headings are descriptive. Support documentation headings have been confirmed. |
| 2.4.7 Focus Visible | Partially Supports | The product is primarily operated by touch and physical buttons. The interface does not provide a conventional keyboard focus indicator. Visual focus/selection is indicated by a background color change on the active field. |
| 3.1.2 Language of Parts | Supports | The interface supports multiple languages. The user guide separates English and French sections. Document language handling has been confirmed. |
| 3.2.3 Consistent Navigation | Not Applicable | This criterion does not apply to non-web software or non-web documents in this report scope. |
| 3.2.4 Consistent Identification | Not Applicable | This criterion does not apply to non-web software or non-web documents in this report scope. |
| 3.3.3 Error Suggestion | Not Applicable | The device does not include form-like input error correction workflows where error suggestions would apply. Troubleshooting information is provided in support documentation. |
| 3.3.4 Error Prevention (Legal, Financial, Data) | Not Applicable | The product does not perform legal, financial, or user-data transactions covered by this criterion. |

# Revised Section 508 Report

The following tables report conformance with the Revised Section 508 standards for the explorē 12 Kit scope described above.

## Chapter 3: Functional Performance Criteria (FPC)

Notes: Chapter 3 is included to explain functional access where Chapters 4 and 5 do not fully describe user access, especially because the product is a low vision product rather than a blindness product.

| Criteria | Conformance Level | Remarks and Explanations |
| --- | --- | --- |
| 302.1 Without Vision | Does Not Support | explorē 12 is intended for users with limited, but remaining usable vision. Core operation depends on visual access to the screen. The device does not provide speech output, braille output, or a fully nonvisual operating mode. |
| 302.2 With Limited Vision | Supports | The product is designed for users with low vision. It provides magnification, contrast modes, brightness control, line and blinds, preferred zoom settings, near/distance viewing, and external display support. |
| 302.3 Without Perception of Color | Supports | Color is not the only means of access. Core controls are identified by placement, labels, icons, and physical layout. Multiple contrast modes are available. |
| 302.4 Without Hearing | Supports | The product does not require hearing for core operation. Information is presented visually, and the product uses vibration for some feedback. |
| 302.5 With Limited Hearing | Supports | The product does not rely on hearing for core operation. |
| 302.6 Without Speech | Supports | The product does not require speech input for operation. |
| 302.7 With Limited Manipulation | Partially Supports | Core functions can be operated through simple physical controls and touchscreen options. Some use cases require holding the device, pressing physical buttons, using touch-based menu selection, docking or removing the device from the folding stand, or manipulating the stand. |
| 302.8 With Limited Reach and Strength | Partially Supports | The device is portable and can be used on a table or stand. However, lifting, holding, docking, opening, collapsing, or transporting the device and stand may require reach or strength. |
| 302.9 With Limited Language, Cognitive, and Learning Abilities | Partially Supports | The interface is simple and task-based, with limited menus and clear control labels. Some setup, menu navigation, software update steps, and support documentation may require reading and following multi-step instructions. |

## Chapter 4: Hardware

Notes: Scope includes explorē 12 kit hardware, including tablet, physical controls, touchscreen, folding stand, USB-C connection, video output, battery/charging, and status indicators.

| Criteria | Conformance Level | Remarks and Explanations |
| --- | --- | --- |
| 402 Closed Functionality | Heading cell - no response required | Heading cell - no response required |
| 402.1 General | Heading cell - no response required | Heading cell - no response required |
| 402.2 Speech-Output Enabled | Heading cell - no response required | Heading cell - no response required |
| 402.2.1 Information Displayed On-Screen | Does Not Support | The product displays information on-screen and does not provide speech output or braille output. It is intended for users with limited, usable vision, not users without vision. |
| 402.2.2 Transactional Outputs | Not Applicable | The product does not provide transactional outputs such as receipts, tickets, or other transaction records. |
| 402.2.3 Speech Delivery Type and Coordination | Not Applicable | The product does not provide speech output. |
| 402.2.4 User Control | Not Applicable | The product does not provide speech output. |
| 402.2.5 Braille Instructions | Not Applicable | The product does not provide speech output and does not include braille operation instructions. |
| 402.3 Volume | Heading cell - no response required | Heading cell - no response required |
| 402.3.1 Private Listening | Not Applicable | The product does not provide speech output or audio output for private listening. |
| 402.3.2 Non-private Listening | Not Applicable | The product does not provide speech output or audio output for non-private listening. |
| 402.4 Characters on Display Screens | Supports | The product provides a 12-inch display, magnification, menu text size options, high-contrast modes, and brightness control. Menu text can be presented in Small, Medium, or Large and is user adjustable. |
| 402.5 Characters on Variable Message Signs | Not Applicable | The product does not include variable message signs. |
| 403 Biometrics | Heading cell - no response required | Heading cell - no response required |
| 403.1 Biometrics | Not Applicable | The product does not use biometrics. |
| 404 Preservation of Information Provided for Accessibility | Heading cell - no response required | Heading cell - no response required |
| 404.1 Preservation of Information Provided for Accessibility | Not Applicable | The product exports saved images. It does not create or transform content with accessibility metadata that must be preserved, and no future file formats are planned that would include accessibility information. |
| 405 Privacy | Heading cell - no response required | Heading cell - no response required |
| 405.1 Privacy | Not Applicable | The product does not provide speech output or other accessibility functions that create a specific privacy concern under this criterion. |
| 406 Standard Connections | Heading cell - no response required | Heading cell - no response required |
| 406.1 Standard Connections | Supports | The product uses standard USB Type-C power/data connection and a high-resolution video output cable for external display use. |
| 407 Operable Parts | Heading cell - no response required | Heading cell - no response required |
| 407.2 Contrast | Supports | Physical button controls are highly contrasted and visually distinguishable by layout, labeling, and color contrast. |
| 407.3 Input Controls | Heading cell - no response required | Heading cell - no response required |
| 407.3.1 Tactilely Discernible | Supports | Primary physical controls are tactilely discernible by location and layout. Controls are discernible by simple layout and color variation. |
| 407.3.2 Alphabetic Keys | Not Applicable | The product does not include alphabetic keys. |
| 407.3.3 Numeric Keys | Not Applicable | The product does not include numeric keys. |
| 407.4 Key Repeat | Supports | Zoom controls support press-and-hold behavior. Key repeat behavior is also supported for scrolling menus. |
| 407.5 Timed Response | Supports | The device auto-off setting is adjustable to 1 minute, 2 minutes, 5 minutes, 10 minutes, or Never. |
| 407.6 Operation | Partially Supports | Core controls do not appear to require tight grasping, pinching, or twisting. Some use of the stand, docking, folding, or carrying may require two hands or moderate dexterity. |
| 407.7 Tickets, Fare Cards, and Keycards | Not Applicable | The product does not dispense or process tickets, fare cards, or keycards. |
| 407.8 Reach Height and Depth | Heading cell - no response required | Heading cell - no response required |
| 407.8.1 Vertical Reference Plane | Not Applicable | The product is portable/personal-use ICT and is not installed in a fixed public location requiring fixed reach-range evaluation. |
| 407.8.1.1 Vertical Plane for Side Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.1.2 Vertical Plane for Forward Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.2 Side Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.2.1 Unobstructed Side Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.2.2 Obstructed Side Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.3 Forward Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.3.1 Unobstructed Forward Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.3.2 Obstructed Forward Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.3.2.1 Operable Part Height for ICT with Obstructed Forward Reach | Not Applicable | Not fixed-location ICT. |
| 407.8.3.2.2 Knee and Toe Space under ICT with Obstructed Forward Reach | Not Applicable | Not fixed-location ICT. |
| 408 Display Screens | Heading cell - no response required | Heading cell - no response required |
| 408.2 Visibility | Supports | The product provides a 12-inch screen, screen brightness control, contrast modes, magnification, and external display support. |
| 408.3 Flashing | Supports | No flashing content occurs during normal operation. |
| 409 Status Indicators | Heading cell - no response required | Heading cell - no response required |
| 409.1 Status Indicators | Partially Supports | Status indicators are visual. Vibration is available for some functions, but battery/status indicators appear visual only. |
| 410 Color Coding | Heading cell - no response required | Heading cell - no response required |
| 410.1 Color Coding | Supports | Color is not the only means used to identify or operate core functions. |
| 411 Audible Signals | Heading cell - no response required | Heading cell - no response required |
| 411.1 Audible Signals | Not Applicable | The product does not rely on audible signals. It uses visual information and vibration for some feedback. |
| 412 ICT with Two-Way Voice Communication | Heading cell - no response required | Heading cell - no response required |
| 412.2 Volume Gain | Heading cell - no response required | Heading cell - no response required |
| 412.2.1 Volume Gain for Wireline Telephones | Not Applicable | The product is not a wireline telephone. |
| 412.2.2 Volume Gain for Non-Wireline ICT | Not Applicable | The product is not non-wireline ICT with two-way voice communication. |
| 412.3 Interference Reduction and Magnetic Coupling | Heading cell - no response required | Heading cell - no response required |
| 412.3.1 Wireless Handsets | Not Applicable | The product is not a wireless handset. |
| 412.3.2 Wireline Handsets | Not Applicable | The product is not a wireline handset. |
| 412.4 Digital Encoding of Speech | Not Applicable | The product does not provide two-way voice communication. |
| 412.5 Real-Time Text Functionality | Reserved for future | Reserved for future |
| 412.6 Caller ID | Not Applicable | The product does not provide caller ID. |
| 412.7 Video Communication | Not Applicable | The product does not provide real-time video communication between users. |
| 412.8 Legacy TTY Support | Heading cell - no response required | Heading cell - no response required |
| 412.8.1 TTY Connectability | Not Applicable | The product is not ICT with voice communication. |
| 412.8.2 Voice and Hearing Carry Over | Not Applicable | The product is not ICT with voice communication. |
| 412.8.3 Signal Compatibility | Not Applicable | The product is not ICT with voice communication. |
| 412.8.4 Voice Mail and Other Messaging Systems | Not Applicable | The product does not provide voice mail or messaging systems. |
| 413 Closed Caption Processing Technologies | Heading cell - no response required | Heading cell - no response required |
| 413.1.1 Decoding and Display of Closed Captions | Not Applicable | The product does not receive, play, or display captioned video programming. |
| 413.1.2 Pass-Through of Closed Caption Data | Not Applicable | The product does not process closed caption data. |
| 414 Audio Description Processing Technologies | Heading cell - no response required | Heading cell - no response required |
| 414.1.1 Digital Television Tuners | Not Applicable | The product does not include a digital television tuner. |
| 414.1.2 Other ICT | Not Applicable | The product does not process audio description content. |
| 415 User Controls for Captions and Audio Descriptions | Heading cell - no response required | Heading cell - no response required |
| 415.1.1 Caption Controls | Not Applicable | The product does not display captions or captioned video programming. |
| 415.1.2 Audio Description Controls | Not Applicable | The product does not play audio description content. |

## Chapter 5: Software

Notes: Scope includes embedded non-web software running on explorē 12 software version 1.1.6.450.

| Criteria | Conformance Level | Remarks and Explanations |
| --- | --- | --- |
| 501.1 Scope - Incorporation of WCAG 2.0 AA | See WCAG 2.0 matrix | The embedded non-web software is addressed in the WCAG 2.0 Level A and Level AA tables above. |
| 502 Interoperability with Assistive Technology | Heading cell - no response required | Heading cell - no response required |
| 502.2.1 User Control of Accessibility Features | Not Applicable | The product is closed-functionality embedded ICT and does not run user-installed assistive technology or platform accessibility features. Accessibility-related settings are provided within the product itself. |
| 502.2.2 No Disruption of Accessibility Features | Not Applicable | The product does not run on a general-purpose platform with separate assistive technology features. |
| 502.3 Accessibility Services | Heading cell - no response required | Heading cell - no response required |
| 502.3.1 Object Information | Not Applicable | The embedded software does not expose object information to platform accessibility services. |
| 502.3.2 Modification of Object Information | Not Applicable | No platform accessibility services are exposed. |
| 502.3.3 Row, Column, and Headers | Not Applicable | The on-device software does not expose tabular object information to platform accessibility services. |
| 502.3.4 Values | Not Applicable | The embedded software does not expose values to platform accessibility services. |
| 502.3.5 Modification of Values | Not Applicable | No platform accessibility services are exposed. |
| 502.3.6 Label Relationships | Not Applicable | No platform accessibility services are exposed. |
| 502.3.7 Hierarchical Relationships | Not Applicable | No platform accessibility services are exposed. |
| 502.3.8 Text | Not Applicable | No platform accessibility services are exposed. |
| 502.3.9 Modification of Text | Not Applicable | The product does not provide user text editing through platform accessibility services. |
| 502.3.10 List of Actions | Not Applicable | No platform accessibility services are exposed. |
| 502.3.11 Actions on Objects | Not Applicable | No platform accessibility services are exposed. |
| 502.3.12 Focus Cursor | Not Applicable | No platform accessibility services are exposed. |
| 502.3.13 Modification of Focus Cursor | Not Applicable | No platform accessibility services are exposed. |
| 502.3.14 Event Notification | Not Applicable | No platform accessibility services are exposed. |
| 502.4 Platform Accessibility Features | Not Applicable | The product does not run on a general-purpose platform with external platform accessibility features. |
| 503 Applications | Heading cell - no response required | Heading cell - no response required |
| 503.2 User Preferences | Supports | The product provides its own user preferences for contrast, menu color, brightness, menu text size, preferred zoom, touch gestures, lights, language, and auto-off behavior. |
| 503.3 Alternative User Interfaces | Not Applicable | No platform accessibility services are exposed. The product provides physical controls and touchscreen controls for core functions. All touchscreen functions are supported through physical controls except selection of items in settings menus, where touch is the only method of selection. |
| 503.4 User Controls for Captions and Audio Description | Heading cell - no response required | Heading cell - no response required |
| 503.4.1 Caption Controls | Not Applicable | The product does not display captioned video. |
| 503.4.2 Audio Description Controls | Not Applicable | The product does not provide audio description playback. |
| 504 Authoring Tools | Heading cell - no response required | Heading cell - no response required |
| 504.2 Content Creation or Editing | Not Applicable | The product is not an authoring tool. It can save images to a gallery, but it does not create/edit content intended for publication with accessibility metadata. |
| 504.2.1 Preservation of Information Provided for Accessibility in Format Conversion | Not Applicable | The product does not perform format conversions of authored content with accessibility information. |
| 504.2.2 PDF Export | Not Applicable | The product does not export PDF files. |
| 504.3 Prompts | Not Applicable | The product is not an authoring tool. |
| 504.4 Templates | Not Applicable | The product is not an authoring tool and does not provide templates. |

## Chapter 6: Support Documentation and Services

Notes: Scope includes user guide, quick start, electronic support files, HumanWare-produced support videos linked from the support page to YouTube, and HumanWare support services.

| Criteria | Conformance Level | Remarks and Explanations |
| --- | --- | --- |
| 601.1 Scope | Heading cell - no response required | Heading cell - no response required |
| 602 Support Documentation | Heading cell - no response required | Heading cell - no response required |
| 602.2 Accessibility and Compatibility Features | Supports | HumanWare can provide information about accessibility and compatibility features. Product documentation describes magnification, contrast, lines/blinds, lights, menu text size, touchscreen configuration, languages, USB export, and external display use. |
| 602.3 Electronic Support Documentation | Supports | Electronic support documentation is available as accessible PDF and Word files. HumanWare-produced support videos linked from the HumanWare support page to YouTube are included in this report and are captioned. See the WCAG 2.0 matrix for criteria-level notes related to electronic documents and videos. |
| 602.4 Alternate Formats for Non-Electronic Support Documentation | Does Not Support | Alternate-format documentation upon request is not currently provided, based on product team input. |
| 603 Support Services | Heading cell - no response required | Heading cell - no response required |
| 603.2 Information on Accessibility and Compatibility Features | Supports | Support services can provide information about accessibility and compatibility features. |
| 603.3 Accommodation of Communication Needs | Supports | Support is available through communication methods that can accommodate disability-related needs, based on product team input. |

# Legal Disclaimer (HumanWare)

This Accessibility Conformance Report is provided for informational purposes only and reflects HumanWare’s evaluation of the product identified above as of the report date and product version stated in this document. The information in this report is based on HumanWare’s review of product functionality, support documentation, support services, and available test information. Product features, documentation, support materials, and services may change without notice. This report is not intended to create any warranty, certification, contractual commitment, or legal obligation beyond those expressly provided in HumanWare’s applicable product documentation, agreements, or warranty terms. Customers should evaluate the product in their own environment to determine whether it meets their accessibility, procurement, and user requirements.

“Voluntary Product Accessibility Template” and “VPAT” are registered service marks of the Information Technology Industry Council (ITI).